Excerpt from: TASCC Code of Practice for Combinable Crops and Animal Feeds effective from October 2024 (Version 6 - 01/07/2026)

# Appendix 17 - Standard Values for Testing

## The standard

This publication describes the elements needed for best laboratory practice for grain testing. It gives straightforward guidance on the requirements for equipment, methods and operators. It is not designed to supersede any quality systems already in place and is subject to third party audit and accreditation. However, it provides a framework for those who do not operate to these standards (see page 2).

In addition to the framework, the publication describes targets for testing. These are measures of the proficiency of a laboratory and are given as the limits of variation that are acceptable for repeat testing by the laboratory. They are also given as limits of deviation from defined mean/median values or reference values when participating in proficiency test schemes and limits of variation for use with Internal Quality Control checks. Limits of variation for staff competency are also provided. These values have been derived from a very large pool of actual testing data. An explanation of terminology and derivation of standard value is given on page 3, followed by standard values in the table on page 4 and onwards.

## Who should adopt this standard?

All UK laboratories that test grain for trading purposes. Registers of participating laboratories (along with their testing protocols) can be found the relevant trade association websites.

Click on the relevant link to view the register or scheme.

Visit the UK Flour Millers site: ukflourmillers.org/flour-milling/the-milling-process/wheat-testing-at-mill-intake

See the Trade Assurance Schemes page on the AIC site: agindustries.org.uk/sectors/trade-assurance-schemes.html

Learn about MBC Protocols and Procedures on the The Maltsters’ Association of Great Britain site: ukmalt.com/uk-malting-industry/malting-barley-committee/mbc-protocols-procedures/

Go to The United Kingdom Accreditation Service (UKAS) site: ukas.com/find-an-organisation/

## Scope – tests to be included in the scheme

Ideally all participating laboratories should include the following tests within ring testing schemes. Those laboratories not undertaking the full range of tests will have a reduced scope.

| Test | Wheat | Barley | OSR | Oats | Rye | Peas | Beans | Linseed | Maize |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Moisture | Included | Included | Included | Included | Included | Included | Included | Included | Included |
| Specific weight | Included | Included | Not included | Included | Included | Not included | Not included | Not included | Not included |
| Screenings | Included | Included | Not included | Included | Included | Not included | Not included | Not included | Not included |
| Admixture | Included | Included | Included | Included | Included | Included | Included | Included | Included |
| Protein | Included | Not included | Not included | Included | Included | Not included | Included | Not included | Not included |
| Nitrogen | Not included | Included | Not included | Not included | Not included | Not included | Not included | Not included | Not included |
| Oil | Not included | Not included | Included | Not included | Not included | Not included | Not included | Included |  |
| Falling number | Included | Not included | Not included | Not included | Included | Not included | Not included | Not included | Not included |
| Gluten | Included | Not included | Not included | Not included | Not included | Not included | Not included | Not included | Not included |
| Hardness | Included | Not included | Not included | Not included | Not included | Not included | Not included | Not included | Not included |
| Germination | Not included | Included | Not included | Not included | Not included | Not included | Not included | Not included | Not included |
| Mycotoxins | Included | Included | Not included | Included | Included | Not included | Not included | Not included | Not included |
| Erucic Acid | Not included | Not included | Included | Not included | Not included | Not included | Not included | Not included | Not included |

## Grain testing best practice

### Equipment

- All equipment used for testing must be fit for purpose, serviced and maintained in compliance with the laboratory’s quality system and manufacturer’s instructions. Requirements for equipment, eg balance accuracy, would be given in the test method

- Equipment shall be calibrated according to routines within the method and/or against a defined schedule

- Performance checks must be done before use. These checks should be performed daily for contractual testing or weekly, as a minimum, for non-contractual testing

- When all the above (1–3) have been fulfilled the equipment is ready for use

### Method

- The method shall be documented – this can be an official method or an in-house procedure. If the latter is used, the relationship to the official method must be known. Procedures must be in place to ensure that only current methods are used

- Is the method accepted by the grain industry? Ensure you only use methods that give results that others recognise, e.g. SKCS scale for wheat hardness

### Operator

- All staff must be competent in those tests that they are expected to undertake

- All staff (full-time and temporary) shall be trained in the methods they are required to perform. A record of competence will be held by the laboratory manager and reviewed regularly. Preferably, staff will have undertaken the necessary laboratory training modules (available at Grain Analyst Training modules (AHDB): ahdb.talentlms.com/plus/login)

- All staff must ensure that any instrument of test performance checks are completed before routine testing can begin. These checks should be performed daily for contractual testing or weekly, as a minimum, for non-contractual testing

### Sample, testing and reporting

- Only when all the above elements have been fulfilled can the sample be taken and testing start. Ensure that the test sample is representative of the sample provided to the laboratory (either in bag or from the laboratory taken sample). Rules about sub-sampling, grinding and mixing must be followed

- Undertake the test according to the documented method. Observe all the critical control points

- Results shall be reported in a format agreed with the customer. All results shall be reported with units of measurement as recognised in the grain industry

### Standard values for testing

- Repeatability – permitted difference between two results on same sample done by the same person, using same instrument within a short timeframe

- Reproducibility – permitted difference between two results on the same sample from different testing facilities

- Standard Deviation (SD) for Proficiency Testing Schemes (PT) – maximum permitted value to be used to calculate z-scores (see page 4 for explanation of z-scores). Providers are encouraged to use lower values for SD if participants are regularly achieving better performance than the standard set by these values

- Standard Deviation (SD) for Internal Quality Control (IQC) – maximum suggested value to be used to calculate monitoring (±1SD), warning (±2SD) and action limits (±3SD) for IQC sample

- Staff competency – Dependent on methodology, a combination of repeatability (column 1) and reproducibility (column 2) values for the test are used. Reproducibility can be checked by labs being members of a Proficiency “ring” check scheme. There are two main steps for checking staff competency:

  - Repeatability – The first step is by repeat testing one sample a minimum of five times, ensuring all results fall within the repeatability value (column 1)

  - Competency – The second step is by testing a minimum of five different samples, representing a range of values, and comparing the results obtained with those of a trained operator/established value for the sample. Difference between results should not exceed the figure given in column 5

Please refer to the Grain Analyst Training modules available on the Grain Analyst Training modules (AHDB) site: ahdb.talentlms.com/plus/login

### Interpretation of z-score

Regardless of the sign (+/-):

- |Z| ≤ 2: performance is satisfactory

- 2 < |Z| ≤ 3: is a warning (occurs 1 in 20) and suggest attention to equipment or procedures may be required

- |Z| > 3: requires immediate action and requires attention to equipment, procedure or operator

The figures presented in the following tables are based on proficiency testing schemes data provided by end-users where significant amounts of data, often a year’s worth, was available. Reference has also been made to standard deviation data published in IOB, EBC, CCAT, BSI and ISO standards.

For a number of tests compromises have had to be made and these are based on the knowledge and experience of members of the Grain Sampling and Analysis Laboratory Working Party. This reflects the fact that the data available from different sources covered different ranges of test results.

#### Wheat

| Commodity / Test | 1 Repeatability (r) | 2 Reproducibility (R) | 3 SD for PT | 4 SD for IQC | 5 Competency |
| --- | --- | --- | --- | --- | --- |
| Moisture |  |  |  |  |  |
| - Oven | 0.20 | 0.36 | 0.20 | N/A | 0.36 |
| - NIR | 0.2 | 0.4 | 0.20 | 0.1 | 0.2 |
| - Meter | 0.3 | 0.5 | 0.30 | 0.15 | 0.3 |
| Specific weight |  |  |  |  |  |
| - Kern | 0.40 | 0.90 | 0.40 | 0.20 | 0.70 |
| - Other | 0.5 | 1.5 | 0.70 | 0.25 | 0.5 |
| Screenings | See Table 1 | See Table 1 | See Table 1 | N/A | See Table 1 |
| Falling number | See Table 4 | See Table 4 | See Table 4 | See Table 4 | See Table 4 |
| Gluten | 1.1 | 2.8 | 1.0 | 0.5 | 2.0 |
| Hardness |  |  |  |  |  |
| - SKCS | 3.1 | 6.7 | 2.4 | 1.6 | 3.1 |
| Protein |  |  |  |  |  |
| - Dumas | 0.25 | 0.40 | 0.14 | 0.125 | 0.35 |
| - NIR (Transmission) | 0.25 | 0.48 | 0.20 | 0.125 | 0.25 |
| - NIR (Reflectance) | 0.30 | 0.60 | 0.20 | 0.15 | 0.50 |
| Admixture | See Table 1 | See Table 1 | See Table 1 | N/A | See Table 1 |
| Mycotoxins | No current data available |  |  |  |  |

#### Barley

| Commodity / Test | 1 Repeatability (r) | 2 Reproducibility (R) | 3 SD for PT | 4 SD for IQC | 5 Competency |
| --- | --- | --- | --- | --- | --- |
| Moisture |  |  |  |  |  |
| - Oven | 0.20 | 0.45 | 0.20 | N/A | 0.45 |
| - NIR | 0.2 | 0.4 | 0.20 | 0.1 | 0.2 |
| - Meter | 0.3 | 0.5 | 0.30 | 0.15 | 0.3 |
| Specific weight |  |  |  |  |  |
| - Kern | 0.40 | 0.90 | 0.40 | 0.20 | 0.70 |
| - Other | 0.6 | 1.5 | 0.70 | 0.3 | 0.6 |
| Screenings | See Table 1 | See Table 1 | See Table 1 | N/A | See Table 1 |
| Nitrogen |  |  |  |  |  |
| - Dumas | 0.05 | 0.09 | 0.03 | 0.025 | 0.07 |
| - NIR (Transmission) | 0.05 | 0.09 | 0.05 | 0.025 | 0.05 |
| - NIR (Reflectance) | 0.05 | 0.11 | 0.05 | 0.025 | 0.09 |
| Germination capacity |  |  |  |  |  |
| - Rapid | See Table 5 | See Table 5 | See Table 5 | N/A | See Table 5 |
| - Peroxide | See Table 5 | See Table 5 | See Table 5 | N/A | See Table 5 |
| - Energy (BRF) | See Table 6 | See Table 6 | See Table 6 | N/A | See Table 6 |
| Admixture | See Table 1 | See Table 1 | See Table 1 | N/A | See Table 1 |
| Mycotoxins | No current data available |  |  |  |  |

#### Oilseed rape

| Commodity / Test | 1 Repeatability (r) | 2 Reproducibility (R) | 3 SD for PT | 4 SD for IQC | 5 Competency |
| --- | --- | --- | --- | --- | --- |
| Moisture |  |  |  |  |  |
| - Oven | 0.20 | 0.40 | 0.20 | N/A | 0.40 |
| - NIR | 0.3 | 0.5 | 0.30 | 0.15 | 0.3 |
| - Meter | 0.3 | 0.5 | 0.30 | 0.15 | 0.3 |
| Admixture |  |  |  |  |  |
| - Reference method | See Table 2 | See Table 2 | See Table 2 | N/A | See Table 2 |
| - Mechanical method | See Table 2 | See Table 2 | See Table 2 | N/A | See Table 2 |
| - Manual method | See Table 2 | See Table 2 | See Table 2 | N/A | See Table 2 |
| Oil content |  |  |  |  |  |
| - Extraction | 0.27 | 1.54 | 0.54 | N/A | 1.54 |
| - NMR | 0.6 | 1.5 | 0.53 | 0.3 | 0.6 |
| - NIR | 0.6 | 1.2 | 0.50 | 0.3 | 0.6 |
| Erucic acid (NIR) | See Table 8 | See Table 8 | See Table 8 | See Table 8 | See Table 8 |

#### Oats

| Commodity / Test | 1 Repeatability (r) | 2 Reproducibility (R) | 3 SD for PT | 4 SD for IQC | 5 Competency |
| --- | --- | --- | --- | --- | --- |
| Moisture |  |  |  |  |  |
| - Oven | 0.32 | 0.54 | 0.20 | N/A | 0.54 |
| - NIR | 0.2 | 0.4 | 0.20 | 0.15 | 0.2 |
| - Meter | 0.3 | 0.5 | 0.30 | 0.15 | 0.3 |
| Bulk Density |  |  |  |  |  |
| - Kern | 0.50 | 1.80 | 1.0 | 0.25 | 1.10 |
| - Other | 0.9 | 1.9 | 1.0 | 0.45 | 0.9 |
| Screenings | See Table 1 | See Table 1 | See Table 1 | N/A | See Table 1 |
| Admixture | See Table 1 | See Table 1 | See Table 1 | N/A | See Table 1 |
| Protein |  |  |  |  |  |
| - Dumas | 0.35 | 0.83 | 0.30 | 0.15 | 0.54 |
| - NIR (Transmission) | 0.51 | 0.57 | 0.25 | 0.25 | 0.51 |

#### Rye

| Commodity / Test | 1 Repeatability (r) | 2 Reproducibility (R) | 3 SD for PT | 4 SD for IQC | 5 Competency |
| --- | --- | --- | --- | --- | --- |
| Moisture |  |  |  |  |  |
| - Oven | 0.20 | 0.36 | 0.20 | N/A | 0.36 |
| - NIR | 0.2 | 0.4 | 0.20 | 0.1 | 0.2 |
| - Meter | 0.3 | 0.5 | 0.30 | 0.15 | 0.3 |
| Bulk Density |  |  |  |  |  |
| - Kern | 0.40 | 0.90 | 0.40 | 0.20 | 0.70 |
| - Other | 0.5 | 1.5 | 0.70 | 0.25 | 0.5 |
| Falling number | See Table 4 | See Table 4 | See Table 4 | See Table 4 | See Table 4 |
| Protein |  |  |  |  |  |
| - Dumas | 0.25 | 0.40 | 0.14 | 0.125 | 0.35 |
| - NIR (Transmission) | 0.25 | 0.48 | 0.20 | 0.125 | 0.25 |
| - NIR (Reflectance) | 0.30 | 0.60 | 0.20 | 0.15 | 0.50 |
| Screenings | See Table 1 | See Table 1 | See Table 1 | N/A | See Table 1 |
| Admixture | See Table 1 | See Table 1 | See Table 1 | N/A | See Table 1 |

#### Peas

| Commodity / Test | 1 Repeatability (r) | 2 Reproducibility (R) | 3 SD for PT | 4 SD for IQC | 5 Competency |
| --- | --- | --- | --- | --- | --- |
| Moisture |  |  |  |  |  |
| - Oven | 0.30 | 0.50 | 0.20 | N/A | 0.50 |
| - Rapid | 0.3 | 0.8 | 0.30 | 0.15 | 0.3 |
| Waste (and Admixture) | See Table 3 | See Table 3 | See Table 3 | N/A | See Table 3 |
| Cracked seed coats | See Table 3 | See Table 3 | See Table 3 | N/A | See Table 3 |
| Soak test | See Table 3 | See Table 3 | See Table 3 | N/A | See Table 3 |

#### Beans

| Commodity / Test |  | 1 Repeatability (r) | 2 Reproducibility (R) | 3 SD for PT | 4 SD for IQC | 5 Competency |
| --- | --- | --- | --- | --- | --- | --- |
| Moisture |  |  |  |  |  |  |
| - Oven |  | 0.30 | 0.50 | 0.20 | N/A | 0.50 |
| - Rapid |  | 0.3 | 0.8 | 0.30 | 0.15 | 0.3 |
| Admixture |  | See Table 3 | See Table 3 | See Table 3 | N/A | See Table 3 |
| Protein |  |  |  |  |  |
| - Dumas | 0.6 | 0.6 | 0.5 | 0.3 | 0.6 |
| - Rapid | 0.6 | 0.6 | 0.5 | 0.3 | 0.6 |
| Staining and Colour |  | See Table 5 | See Table 5 | See Table 5 | See Table 5 | See Table 5 |
| Insect Test |  | See Table 7 | See Table 7 | See Table 7 | See Table 7 | See Table 7 |

#### Linseed

| Commodity / Test | 1 Repeatability (r) | 2 Reproducibility (R) | 3 SD for PT | 4 SD for IQC | 5 Competency |
| --- | --- | --- | --- | --- | --- |
| Moisture |  |  |  |  |  |
| - Oven | 0.20 | 0.40 | 0.20 | N/A | 0.40 |
| - Rapid | 0.3 | 0.5 | 0.30 | 0.15 | 0.3 |
| Admixture | See Table 2 | See Table 2 | See Table 2 | N/A | See Table 2 |
| Oil content |  |  |  |  |  |
| - Extraction | 0.27 | 1.54 | 0.54 | N/A | 1.54 |
| - Other | 0.6 | 1.5 | 0.54 | 0.3 | 0.6 |

#### Maize

| Commodity / Test | 1 Repeatability (r) | 2 Reproducibility (R) | 3 SD for PT | 4 SD for IQC | 5 Competency |
| --- | --- | --- | --- | --- | --- |
| Moisture |  |  |  |  |  |
| - Oven | 0.30 | 0.60 | 0.20 | N/A | 0.60 |
| - Rapid | 0.3 | 0.8 | 0.3 | 0.15 | 0.3 |
| Admixture | See Table 3 | See Table 3 | See Table 3 | N/A | See Table 3 |

#### Balance checks

(Balances are checked with known value weights in the critical range, upper and lower action lines are set using the figures in the table below)

| Number of decimal places of balance (dp) | Maximum tolerance for the check weight |
| --- | --- |
| 1 dp (0.1g) | +/- 0.3g |
| 2 dp (0.01g) | +/- 0.03g |
| 3 dp (0.001g) | +/- 0.003g |
| 4 dp (0.0001g) | +/- 0.0003g |

#### Dispenser checks

(Dispensers are checked by weighing the volume of dispensed water)

(based on 1ml = 1g)

| Volume of dispenser | Below weights are based on a tolerance of 5% |  |
| --- | --- | --- |
| 100 µl | 0.1g +/- 0.005g |  |
| 200 µl | 0.2g +/-0.01g |  |
| 300 µl | 0.3g +/- 0.015g |  |
| 400 µl | 0.4g +/- 0.02g |  |
| 500 µl | 0.5g +/-0.025g |  |
| 600 µl | 0.6g +/-0.03g |  |
| 700 µl | 0.7g +/- 0.035g |  |
| 1 ml | 1g +/-0.05g |  |
| For dispensers used with 70% Methanol |  |  |
| 10 ml (70% Methanol); 8.72g +/-0.44g. For dispensers used with Falling Number apparatus |
| 25 ml |  |  | 25g +/- 0.20g |  |

Table 1. Wheat, barley, oats, rye screenings and admixture (Including skinned barley and broken grains)

| Result range |  | 1 Repeatability (r) | 2 Reproducibility (R) | 3 SD for PT | 4 SD for IQC | 5 Competency |
| --- | --- | --- | --- | --- | --- | --- |
| 0.0 to 1.0 | 99.0 to 100 | 0.4 | 0.7 | 0.25 | N/A | 0.7 |
| 1.1 to 2.0 | 98.0 to 98.9 | 0.7 | 1.4 | 0.50 | N/A | 1.4 |
| 2.1 to 5.0 | 95.0 to 97.9 | 1.1 | 2.1 | 0.75 | N/A | 2.1 |
| 5.1 to 10.0 | 90.0 to 94.9 | 1.4 | 2.8 | 1.0 | N/A | 2.8 |
| 10.1 to 20.0 | 80.0 to 89.9 | 2.1 | 4.2 | 1.5 | N/A | 4.2 |
| 20.1 to 30.0 | 70.0 to 79.9 | 3.5 | 7.1 | 2.5 | N/A | 7.1 |
| 30.1 to 50.0 | 50.0 to 69.9 | 7.1 | 14.2 | 5.0 | N/A | 14.2 |

Table 2. Oilseeds admixture

| Result range | 1 Repeatability (r) | 2 Reproducibility (R) | 3 SD for PT | 4 SD for IQC | 5 Competency |
| --- | --- | --- | --- | --- | --- |
| 0.0 to 0.5 | 0.2 | 0.4 | 0.14 | N/A | 0.4 |
| 0.6 to 1.0 | 0.4 | 0.8 | 0.28 | N/A | 0.8 |
| 1.1 to 2.0 | 0.6 | 1.8 | 0.64 | N/A | 1.8 |
| 2.1 to 3.0 | 0.8 | 2.4 | 0.85 | N/A | 2.4 |
| 3.1 to 4.0 | 1.0 | 3.0 | 1.06 | N/A | 3.0 |
| 4.1 to 5.0 | 1.2 | 3.6 | 1.27 | N/A | 3.6 |
| 5.1 to 6.0 | 1.4 | 4.2 | 1.48 | N/A | 4.2 |
| Over 6.0 | 1.6 | 4.8 | 1.70 | N/A | 4.8 |

Table 3. Pea, bean and maize admixture (Including pea insect damage and beans waste (admixture and screenings)).

| Result range |  | 1 Repeatability (r) | 2 Reproducibility (R) | 3 SD for PT | 4 SD for IQC | 5 Competency |
| --- | --- | --- | --- | --- | --- | --- |
| 0.0 to 1.0 | 99.0 to 100 | 0.4 | 1.4 | 0.50 | N/A | 1.4 |
| 1.1 to 2.0 | 98.0 to 98.9 | 0.7 | 2.1 | 0.75 | N/A | 2.1 |
| 2.1 to 3.5 | 96.5 to 97.9 | 1.1 | 3.5 | 1.25 | N/A | 3.5 |
| 3.6 to 5.0 | 95.0 to 96.4 | 1.4 | 4.2 | 1.50 | N/A | 4.2 |
| 5.1 to 10.0 | 90.0 to 94.9 | 2.1 | 5.0 | 1.75 | N/A | 5.0 |
| 10.1 to 20.0 | 80.0 to 89.9 | 3.5 | 7.1 | 2.50 | N/A | 7.1 |
| 20.1 to 30.0 | 70.0 to 79.9 | 5.3 | 10.6 | 3.75 | N/A | 10.6 |
| 30.1 to 50.0 | 50.0 to 69.9 | 7.1 | 14.2 | 5.00 | N/A | 14.2 |

Table 4. Falling number

| Result range | 1 Repeatability (r) | 2 Reproducibility (R) | 3 SD for PT | 4 SD for IQC | 5 Competency |
| --- | --- | --- | --- | --- | --- |
| 60 to 79 | 10 | 10 | 3.5 | 5 | 10 |
| 80 to 109 | 13 | 21 | 7.4 | 7 | 17 |
| 110 to 139 | 15 | 30 | 10.6 | 8 | 23 |
| 140 to 169 | 17 | 38 | 13.4 | 9 | 28 |
| 170 to 199 | 19 | 46 | 16.3 | 10 | 33 |
| 200 to 229 | 21 | 54 | 19.1 | 11 | 38 |
| 230 to 259 | 23 | 62 | 21.9 | 12 | 43 |
| 260 to 289 | 25 | 70 | 24.7 | 13 | 48 |
| 290 to 319 | 27 | 78 | 27.6 | 14 | 53 |
| 320 to 349 | 30 | 86 | 30.4 | 15 | 58 |
| 350 to 379 | 32 | 94 | 33.2 | 16 | 63 |
| Over 379 | 40 | 100 | 35.3 | 20 | 70 |

Table 5. Barley germinative capacity – peroxide and rapid beans – Staining and colour

| Result range | 1 Repeatability (r) | 2 Reproducibility (R) | 3 SD for PT | 4 SD for IQC | 5 Competency |
| --- | --- | --- | --- | --- | --- |
| 100 | 1 | 2 | 0.6 | N/A | 2 |
| 99 | 1 | 2 | 0.8 | N/A | 2 |
| 98 | 2 | 3 | 1.0 | N/A | 3 |
| 97 | 2 | 3 | 1.1 | N/A | 3 |
| 96 | 3 | 4 | 1.3 | N/A | 4 |
| 95 | 3 | 4 | 1.5 | N/A | 4 |
| 94 | 4 | 5 | 1.7 | N/A | 5 |
| 93 | 5 | 5 | 1.9 | N/A | 5 |
| 92 | 5 | 6 | 2.0 | N/A | 6 |
| 91 | 6 | 6 | 2.2 | N/A | 6 |
| 90 | 7 | 7 | 2.4 | N/A | 7 |
| 89 | 7 | 7 | 2.6 | N/A | 7 |
| 88 | 8 | 8 | 2.7 | N/A | 8 |
| 87 | 9 | 9 | 2.9 | N/A | 9 |
| 86 | 9 | 9 | 3.1 | N/A | 9 |

Table 6. Barley germinative energy – BRF

| Result range | 1 Repeatability (r) | 2 Reproducibility (R) | 3 SD for PT | 4 SD for IQC | 5 Competency |
| --- | --- | --- | --- | --- | --- |
| 100 | 3 | 3 | 1.1 | N/A | 3 |
| 99 | 3 | 3 | 1.1 | N/A | 3 |
| 98 | 4 | 4 | 1.4 | N/A | 4 |
| 97 | 4 | 5 | 1.8 | N/A | 5 |
| 96 | 5 | 6 | 2.1 | N/A | 6 |
| 95 | 5 | 7 | 2.5 | N/A | 7 |
| 94 | 6 | 7 | 2.5 | N/A | 7 |
| 93 | 6 | 8 | 2.8 | N/A | 8 |
| 92 | 6 | 8 | 2.8 | N/A | 8 |
| 91 | 7 | 9 | 3.2 | N/A | 9 |
| 90 | 7 | 9 | 3.2 | N/A | 9 |
| 89 | 7 | 10 | 3.5 | N/A | 10 |
| 88 | 8 | 10 | 3.5 | N/A | 10 |
| 87 | 8 | 11 | 3.9 | N/A | 11 |
| 86 | 8 | 11 | 3.9 | N/A | 11 |

Table 7. Beans – Insect test (Bruchid Beetle)

| Result | 1 Repeatability (r) | 2 Reproducibility (R) | 3 SD for PT | 4 SD for IQC | 5 Competency |
| --- | --- | --- | --- | --- | --- |
| 0 | 5.0 | 8.6 | 3.1 | N/A | 8.6 |
| 1 | 5.7 | 9.7 | 3.5 | N/A | 9.7 |
| 2 | 6.3 | 10.8 | 3.9 | N/A | 10.8 |
| 3 | 7.0 | 11.9 | 4.3 | N/A | 11.9 |
| 4 | 7.5 | 12.8 | 4.6 | N/A | 12.8 |
| 5 | 8.2 | 13.9 | 5.0 | N/A | 13.9 |
| 6 | 8.8 | 15.0 | 5.4 | N/A | 15.0 |
| 7 | 9.3 | 15.8 | 5.7 | N/A | 15.8 |
| 8 | 9.9 | 16.9 | 6.1 | N/A | 16.9 |
| 9 | 10.6 | 18.0 | 6.5 | N/A | 18.0 |
| 10 | 11.2 | 19.1 | 6.9 | N/A | 19.1 |
| 11 | 11.7 | 20.0 | 7.2 | N/A | 20.0 |
| 12 | 12.4 | 21.1 | 7.6 | N/A | 21.1 |
| 13 | 13.0 | 22.2 | 8.0 | N/A | 22.2 |
| 14 | 13.5 | 23.0 | 8.3 | N/A | 23.0 |
| 15 | 14.1 | 24.1 | 8.7 | N/A | 24.1 |
| 16 | 14.8 | 25.2 | 9.1 | N/A | 25.2 |
| 17 | 15.4 | 26.3 | 9.5 | N/A | 26.3 |
| 18 | 16.0 | 27.2 | 9.8 | N/A | 27.2 |
| 19 | 16.6 | 28.3 | 10.2 | N/A | 28.3 |
| 20 | 17.2 | 29.4 | 10.6 | N/A | 29.4 |
| 21 | 17.7 | 30.2 | 10.9 | N/A | 30.2 |
| 22 | 18.4 | 31.3 | 11.3 | N/A | 31.3 |
| 23 | 19.0 | 32.4 | 11.7 | N/A | 32.4 |
| 24 | 19.5 | 33.3 | 12.0 | N/A | 33.3 |
| ≥25 | 20.2 | 34.4 | 12.4 | N/A | 34.4 |

Table 8. Oilseed Rape – Erucic acid (NIR)

| Result range | 1 Repeatability (r) | 2 Reproducibility (R) | 3 SD for PT | 4 SD for IQC | 5 Competency |
| --- | --- | --- | --- | --- | --- |
| 0.00 - 1.00 | 1.0 | 2.0 | 0.7 | 0.4 | 1.0 |
| 1.01 - 3.00 | 1.5 | 2.5 | 0.9 | 0.5 | 1.5 |
| 3.01 - 6.00 | 2.0 | 3.0 | 1.2 | 0.7 | 2.0 |
| >6.01 | 3.0 | 4.5 | 1.6 | 1.0 | 3.0 |